## Slide 1
Optical Analysis of Liquid Crystal Lens
1 Douk-Hoon Kim∙ Min-Ho Lee2
1Deaprtment of Optometry, Masan University, 2Ofren Ltd.
INTRODUCTION
노안의 교정시력은 방법은 광학적, 의학적, 약리적, 등 다양하게 시행되고 있다.광학적 방법으로는 기능성렌즈의 사용이 보편적으로 이루어지고 있다. 이는 단초점, 이중초점, 다초점 렌즈 등이 사용되고 있다. 이 같은 기능성 렌즈는 초기 노안의 시력 개선에 좋은 영향을 미칠 수 있다. 그러나 소비자의 경제성과 적용에서 어려움을 가진다. 이 같은 문제점을 보완하고자 본 연구에서 개발한 액정렌즈를 광학적 특성을 분석하여 향후 임상에 적용하는데 자료를 획득하고자 한다.
RESULTS
청색광 투과률[380-500nm]측정에서 코팅렌즈는 93.64%이나 코팅하지 않은 렌즈는 91.00%이다. 흐림도 [Haze] 분석에서는 코팅렌즈는 1.91 +/- 0.001%이나, 코팅하지 않은 렌즈는 2.21+/-0.00%이다. 투영도[Clarity] 분석에서는 코팅렌즈와 코팅하지 않은 렌즈는 모두 0.00+/-0.00%이다. 렌즈의 중심두께는 코팅렌즈와 코팅하지 않는 렌즈 모두 2.00mm이다.
발수각 분석에서는 코팅렌즈는 angle L 43.45degree, angle R 48.90 degree, angle mean 46.18degree이나, 코팅 하지 않는 렌즈는 angle L 76.50degree, angle R 76.61 degree, angle mean 76.56degree이다. 자외선 투과률 분석에서 코팅렌즈는 280nm에서 315nm 사이는 0.01%, 315nm에서 380nm사이는 66.71%, 380nm에서 780nm 사이는 94.66%이나, 코팅하지 않는 렌즈는 280nm에서 315nm 사이는 31.27%, 315nm에서 380nm사이는 81.56%, 380nm에서 780nm 사이는 91.39%이다. 렌즈의 굴절력은 코팅렌즈와 코팅하지 않는 렌즈 모두 +1.25diopter이다. 한편 내구성 환경분석에서는 코팅렌즈와 코팅하지 않는 렌즈 모두 이상 없음으로 분석되었다.
METHODS
액정 렌즈의 광학적 특성 분석은 코팅렌즈와 비코팅 렌즈로 구분해서 하였다.
분석항목은 투과률 측정[Dual Beam UV-Vis Spectrometer, USA], 흐림도 측정[Anti-Abrasin Tester, Germany], 투영도 측정[[Anti-Abrasin Tester, Germany], 렌즈 투께 측정[Thickness gauge, Germany], 발수각 측정[Contact Tester, USA], 내구성 측정[Accelerated Weather Testing, USA], 렌즈 굴절력 측정[Lensmeter, Germany]하였다.
CONCLUSIONS
액정렌즈의 광학적 특성에서 코팅렌즈와 코팅하지 않는 렌즈는 약간의 차이를 확인 할 수 있었다. 코팅한 렌즈는 안경렌즈로서의 사용에 적합한 것으로 확인 할 수 있다. 그러나 향후 렌즈 표면의 미세구조와 성분 분석 및 임상적인 연구가 수행되어야 한다고 생각된다.
REFERENCES
2019 한국안광학회 동계학술대회
[1] Cochener BI, Lafuma A, Khoshnood B, Courouve L, Berdeaux G.Comparison of outcomes with multifocal intraocular lenses: a metaanalysis. Clin Ophthalmol. 2011 Jan 7;5:45-56.
[2]. Khandelwal SS1,2, Jun JJ3, Mak S3,4, Booth MS5, Shekelle PG6.Effectiveness of multifocal and monofocal intraocular lenses for cataract surgery and lens replacement: a systematic review and meta-analysis. Graefes Arch Clin Exp Ophthalmol. 2019 May;257(5):863-875.
* 본 연구는 산업통상자원부 소비제 사업과제로 수행된 것임

## Slide 2
Effect of soft contact lenses on the corneal thickness and
anterior chamber depth in the adult population
Douk-Hoon Kim1
1Department of Optometry, Masan University
2019 한국안광학회 동계학술대회
INTRODUCTION
This study was perform the effect of soft contact lenses on the corneal thickness and anterior chamber depth of the adult Korean Population..
METHODS
The subjects consisted of 102 eyes (31 females, 31 males) with the mean age of 20.97±3.21 years (range 20-23 years) in female and of 23.61±7.52 years (range 20-25 years) in male during 2018. The thinnest cornea thickness(TCT)s , central cornea thickness(CCT) , and pupil centre thickness(PCT) of the corneal thickness were measured using the Pentacam pachymetry. ACD value was also measured with Pentacam.
RESULTS
There was a statistically significant among the TCT, CCT, and PCT values(p<0.001) in female and male. However, there was no statistical difference (p>0.5) in the mean TCT, ACT, PCT between female and male. Also, correlation between ACD and corneal thickness in all subjects had no statistical differences (p>0.5) in all subjects. The analysis of female and male of the thinnest zone, central cornea, pupil center in the corneal thickness and anterior chamber depth for these subjects were summarized in the table 1
Table1. The distribution of the corneal thicknesses and ACD values in adult Korean subjects.
_________________________________________________________________________
Items Female male r p-values . __________________________________________________________________________
Ages 20.967±3.218um 23.606±7.523um
TCT 543.941±56.081μm 527.255±94.309μm .029 .848
CCT 552.392±50.887μm 556.784±33.557μm 0.93 .516
PCT 554.353±51.501μm 558.588±32.376μm .029 .346
ACD 3.024±0.283 mm 3.241±0.344mm .-066 .643
__________________________________________________________________________
TCT: thinnest area of cornea, CCT: central area of cornea, PCT: pupil area of cornea,
ACD: anterior chamber depth. p<0.05*, p<0.01**
CONCLUSIONS
These results suggested that the regional thickness of cornea and ACD with Pentacam can provide correct and useful diagnostic information of the morphology of the RGP contact lens fitting and diagnosis of corneal refraction surgery.
REFERENCES
[1]. Uçakhan ÖÖ, Özkan M, Kanpolat A, 2006. Corneal thickness measurements in normal and keratoconic eyes: Pentacam comprehensive eye scanner versus noncontact specular microscopy and ultrasound pachymetry. J Cataract Refract Surg, 32(6):970-977.
[2]. Ponce CMP, Rocha KM, Smith SD, et al., 2009. Central and peripheral corneal thickness measured with optical coherence tomography, Scheimpflug imaging, and ultrasound pachymetry in normal, keratoconus-suspect, and post-laser in situ keratomileusis eyes. J Cataract Refract Surg, 35(6):1055-1062.

## Slide 3
INTRODUCTION
굴절이상 눈에 대한 교정시력의 개선 방법은 광학적, 의학적, 약리적, 기능적 등 다양하게 시행되고 있다. 현재 가장 널리 이용되고 있는 방법은 광학적 방법으로 안경 및 콘택트렌즈 등이 사용되고 있다. 의료산업의 발달로 시기능 훈련을 통한 시력 개선이 선진국에서 수행되고 있다. 본 연구에서 개발한 시기능 훈련렌즈의 광학적 특성을 분석하여 향후 임상에 적용하는데 기초 자료를 획득하고자 한다.
RESULTS
시기능 훈련렌즈 [Optotune] 청색광 투과률[380-500nm]dms 94.20%이다. 흐림도 [Haze] 분석에서는 48.73+/-0.641%이며, 투영도[Clarity] 분석에서는 0.00+/-0.00%이다. 렌즈의 중심두께는 8.50mm이다. 발수각 분석에서는 angle L 47.54degree, angle R 47.23 degree, angle mean 46.89degree이다. 자외선 투과률 분석에서 280nm에서 315nm 사이는 4.661%, 315nm에서 380nm사이는 38.15%, 380nm에서 780nm 사이는 95.60%이다. 굴절력은 -15.25diopter이다. 한편 내구성 환경분석에서는 이상 없음으로 분석되었다.
METHODS
시기능 훈련 렌즈의 광학적 특성 분석 사용한 방법은 투과률 측정[Dual Beam UV-Vis Spectrometer, USA], 흐림도 측정[Anti-Abrasin Tester, Germany], 투영도 측정[[Anti-Abrasin Tester, Germany], 렌즈 투께 측정[Thickness gauge, Germany], 발수각 측정[Contact Tester, USA], 내구성 측정[Accelerated Weather Testing, USA]를 사용하였다.
CONCLUSIONS
시기능 훈련 렌즈의 광학적 특성은 교정시력 렌즈로서 사용에 적합한 것으로 확인 할 수 있다. 그러나 향후 렌즈 표면의 미세구조와 성분 분석을 통해서 검증을 받고, 임상적인 자료의 추가 연구가 필요한 것으로 여겨진다.
REFERENCES
2018 한국안광학회 동계학술대회
Optical Analysis of Visual Training Lens
Douk-Hoon Kim1 ∙ Soung-Jong Park2
1Deaprtment of Optometry, Masan University, 2,Eden Lux Ltd.
[1] Cochener BI, Lafuma A, Khoshnood B, Courouve L, Berdeaux G.Comparison of outcomes with multifocal intraocular lenses: a metaanalysis. Clin Ophthalmol. 2011 Jan 7;5:45-56.
[2]. Khandelwal SS1,2, Jun JJ3, Mak S3,4, Booth MS5, Shekelle PG6.Effectiveness of multifocal and monofocal intraocular lenses for cataract surgery and lens replacement: a systematic review and meta-analysis. Graefes Arch Clin Exp Ophthalmol. 2019 May;257(5):863-875.
* 본 연구는 산업통상자원부 사업화 과제로 수행된 것임

## Slide 4
Analysis on the corneal thickness of the
young children in Korean Male
Douk-Hoon Kim1
1Department of Optometry, Masan University
2019 한국안광학회 동계학술대회
INTRODUCTION
This study was to analysis the corneal thickness of 14-year-old Korean male population with corneal topography .
METHODS
The subjects consisted of 48 eyes with the 14 year old Korean male during 2018. The thinnest cornea thickness(TCT), central cornea thickness(CCT), pupil centre thickness(PCT), and periphery corneal thickness was measured using the Pentacam pachymetry.
RESULTS
There was no statistically significant in the mean among the TCT, CCT, and PCT , and periphery corneal thickness values(p>0.1) between right eye and left eye. However, there was a statistical significant (p<0.01) in the mean among the TCT, CCT, and PCT , and periphery corneal thickness values(p>0.1) in the right eye and left eye. Also, Korean male of 14-year old have markedly thicker corneas than the India, Chinese, and Taiwan populations. There was no significant difference in the mean thickness between right eye and left eye. The analysis of right eye and left eye of the thinnest zone, central cornea, pupil center in the corneal thickness and anterior chamber depth for these subjects were summarized in the table 1.
CONCLUSIONS
Korean male of 14-year old have markedly thicker corneas than another country populations. These results suggested that the regional thickness data of cornea with Pentacam tool can provide correct and useful diagnostic information of the corneal shape of the contact lens fitting and diagnosis of corneal refraction surgery..
REFERENCES
[1] Al-Farhan HM, Al-Otaibi WM. Comparison of central corneal thickness measurements using ultrasound pachymetry, ultrasound biomicroscopy, and the Artemis-2 VHF scanner in normal eyes. Clin Ophthalmol. 2012, 6:1037–104.
[2]. Williams R et al. Central corneal thickness measurements: using an ultrasonic instrument and 4 optical instruments. Cornea. 2011, 30(11):1238–1243..
Items | OD | OS | t | p-values
PCT | 559.917±27.593μm | 563.500±33.630μm | -.430 | .671
TCT | 553.042±29.359μm | 553.125±32.562μm | -.009 | .993
CCT | 554.917±29.074μm | 561.875±30.837μm | -.831 | .414
MCT | 579.417±30.326μm | 586.792±35.024μm | -.898 | .379
SCT | 567.083±27.921μm | 575.875±32.255μm | -1.086 | .289
LCT | 557.375±27.846μm | 558.125±33.671μm | .-.084 | .934
ICT | 567.542±28.927μm | 567.125±32.847μm | .053 | .958
TCT: thinnest area of cornea thickness, CCT: central cornea thickness, PCT: pupil area of cornea thickness, MCT: medial corneal thickness, SCT: superior corneal thickness, LCT: lateral corneal thickness, ICT: inferior corneal thickness |  |  |  | 
Table 1. The mean thickness of the cornea in 14 year old Korean male subjects. (n=24)